## Slide 1: Tutorial sesion: Mueller Matrix Ellipsometry
Oriol Arteaga
Dep. Applied Physics and Optics
University of Barcelona

## Slide 2
Outline
Historical introduction
Basic concepts about Mueller matrices
Mueller matrix ellipsometry instrumentation
Further insights. Measurements and simulations
Symmetries and asymmetries of the Mueller matrix. Relation to anisotropy.
Applications and examples
Concluding remarks

## Slide 3
Historical introduction

## Slide 4
Historical introduction
G G. Stokes
in 1852
Stokes Parameters
Francis Perrin
in 1942
90 years, almost forgotten!
F. Perrin, J. Chem. Phys. 10, 415 (1942). Translation from the french:
F. Perrin, J. Phys. Rad. 3, 41 (1942)

## Slide 5
1852
G G. Stokes (1819-1903)
Stokes Parameters
1942
Francis Perrin (1901-1992)
1929 Paul Soleillet (1902-1992)
1943
Hans Mueller (1900-1965)
K. Järrendahl and B. Kahr, Woollam newsletter, February 2011, pp. 8–9
F. Perrin, J. Phys. Rad. 3, 41 (1942)
P. Soleillet, Ann. Phys. 12, 23 (1929)
H. Mueller, Report no. 2 of OSR project OEMsr-576 (1943)
Historical introduction

## Slide 6
P. S. Hauge, Opt. Commun. 17, 74 (1976).
R. M. A. Azzam, Opt. Lett. 2, 148-150 (1978).
Instrumental papers about the dual rotating compensator technique
“Mueller matrix ellipsometry”
Web of Science Citation Reports
“Generalized ellipsometry”
Historical introduction
“Mueller matrix spectroscopic ellipsometry”

## Slide 7
Basic concepts about Mueller matrices

## Slide 8: Basic concepts about Mueller matrices
I Intensity
p Degree of polarization
χ Azimuth
φ Ellipticity
No depolarization:
Phenomenological description of any scattering experiment

## Slide 9: Basic concepts about Mueller matrices. No depolarization
A nondepolarizing Mueller matrix is called a Mueller-Jones matrix
Equivalence
is 2x2 complex matrix
M is 4x4 real
Transformation
A Jones or Mueller-Jones Jones depends on 6-7 parameters.
But note that the 16 elements of a Mueller-Jones matrix can be still all different!

## Slide 10: Basic concepts about Mueller matrices. No depolarization and isotropy
All modern ellipsometers measure elements of the Mueller matrix.
This is a common representation for isotropic media:
p
s
p
s
This Mueller matrix depends only on 2 parameters
Standard ellipsometry:
Thickness measurements of thin films
Optical functions of isotropic materials

## Slide 11: Basic concepts about Mueller matrices. Depolarization
Depolarization is the reduction of the degree of polarization of light. Typically occurs when the emerging light is composed of several incoherent contributions.
Quantification of the depolarization: Depolarization index (DI)
The DI of a Mueller-Jones matrix is 1
Reasons:
J. J. Gil, E. Bernabeu, Opt. Acta 32 (1985) 259
Sample exhibits spatial, temporal or frequency heterogeneity over the illuminated area

## Slide 12
Mueller matrix ellipsometry instrumentation

## Slide 13: Mueller matrix ellipsometry instrumentation
PSG
PSA
Polarization state generator: PSG
Polarization state analyzer: PSA
In a MM ellipsometer the PSG and PSA typically contain:
A polarizer (P)
A compensating or retarding element (C)
One exception: division-of-amplitude ellipsometers
P
C
P
C

## Slide 14: Mueller matrix ellipsometry instrumentation
The compensating element is the main difference between different types of Mueller matrix ellipsometers
Rotating Retarders
Fixed Retardation
Changing azimuth
Liquid cristal cells
Waveplates are not very acromatic
Fresnel rohms are hard to rotate
Mechanical rotation
Not transparent in the UV
Temperature dependence
No frequency domain analysis
Variable Retardation (nematic LC)
Changing azimuth (ferroelectric LC)
Piezo-optic modulators
(photoelastic modulators)
Electro-optic modulators
(Pockels cells)
Two PEMs for each PSG or PSA
Too fast for imaging
Variable Retardation
Fixed azimuth
Variable Retardation
Fixed azimuth
Two cells for each PSG or PSA
Small acceptance angle
Too fast for imaging
P. S. Hauge, J. Opt. Soc. Am. 68, 1519-1528 (1978)
E. Garcia-Caurel et al. Thin Solid Films 455 120-123 (2004).
O. Arteaga et al. Appl. Optics 51.28 6805-6817 (2012).
R. C. Thompson et al. Appl. Opt. 19, 1323–1332 (1980).

## Slide 15: Mueller matrix ellipsometry instrumentation
The PSA and PSG of Mueller matrix ellipsometers are no different from other Mueller matrix polarimetric approaches
Mueller matrix microscope with two rotating compensators
spectroscopic polarimeter based on four photoelastic modulators
Normal-incidence reflection imaging based on liquid crystals
Instrumentally wise no different from a MM ellipsometer. Lots of imaging applications in chemistry, medicine, biology, geology, etc.
O. Arteaga et al, Appl. Opt. 53, 2236-2245 (2014)
80 um

## Slide 16
Further insights
Measurement and simulations

## Slide 17: Further insights. Measurement and simulations
A spectroscopic Mueller matrix ellipsometer produces this type of data:
Is this MM depolarizing?
If the depolarization is not significative we can find a proper non-depolarizing estimate

## Slide 18: Further insights. Measurement and simulations
Measurement:
Mueller matrix
Simulation usually generates a
Jones/Mueller-Jones matrix
(coherent model)
Objective: Finding a good nondepolarizing estimate (a Mueller-Jones matrix) for a experimental Mueller matrix
One option,
Cloude estimate using the Cloude sum decomposition
S. R. Cloude, Optik 75, 26 (1986).
R. Ossikovski, Opt. Lett. 37, 578-580 (2012).

## Slide 19: Further insights. Measurement and simulations. Example
Experimental Mueller matrix
1. Calculate the Coherency matrix
Coherency matrix, H
2. Calculate the eigenvectors of H (is a hermitian matrix, so eigenvectors are real)

## Slide 20: Further insights. Measurement and simulations. Example
Best nondepolarizing estimate
Initial Experimental Mueller matrix
Suitable to compare with coherent models
3. The eigenvector corresponding to defines the Jones matrix corresponding to
Jones matrix

## Slide 21
1
2
3
This notation is very suitable for normal-incidence transmission and reflection data:
CD: circular dichroism/diatt.
CB: circular birefrigence/retard.
LD: horiz. linear dichroism/diatt.
... etc
For the previous example:
1
3
2
Further insights. Measurement and simulations. Expressing non-depolarizing data
O. Arteaga & A. Canillas, Opt. Lett. 35, 559-561 (2010)

## Slide 22
Mueller matrix symmetries and anisotropy

## Slide 23: Mueller matrix symmetries and anisotropy
In the isotropic case
The MM elements with an asteriks vanish in absence of absorption and J is real (asumming semi-infinite substrate as a sample)
But this symmetry also applies to some situations with anisotropy!
1
Uniaxial
Biaxial (orthorombic)
Biaxial (monoclinic)
Arrow is P. A.
Arrows are O. A.

## Slide 24: Mueller matrix symmetries and anisotropy
The MM elements with an asteriks vanish in absence of absorption and J is real (asumming semi-infinite substrate as a sample)
2
Uniaxial
Biaxial (orthorombic)
Biaxial (monoclinic)
Arrow is P. A.
Arrows are O. A.

## Slide 25: Mueller matrix symmetries and anisotropy
The MM elements with an asteriks vanish in absence of absorption and J is real (asumming semi-infinite substrate as a sample)
3
Uniaxial
Biaxial (orthorombic)
Biaxial (monoclinic)
Arrow is O. A.
Arrows are O. A.
Arrow is P. A.

## Slide 26: Mueller matrix symmetries and anisotropy
The MM elements with an asteriks vanish in absence of absorption and J is imaginary (asumming semi-infinite substrate as a sample)
4
Bi-isotropic media

## Slide 27
Applications and examples

## Slide 28: Applications and examples. A general idea about anisotropy
Instrinsic anisotropy vs structural/form anisotropy
Expect small values of these elements for intrinsic anisotropy
E.g. Reflection on a calcite substrate
AOI 65o

## Slide 29: Applications. Dielectric tensor of crystals
Measure the complex dielectric function (DF) tensor above and below the band edge
Berreman’s 4x4 complex formalism is used to calculate ρ, ρsp and ρps from elements of and the angle of incidence (a fully analytical treatment is sometimes possible).
The dielectric tensor is symmetric
The principal values of the tensor correspond to crystal symmetry directions for isotropic, uniaxial and orthorhombic materials
A magnetic field breaks the symmetry. E.g. MOKE

## Slide 30: Applications. Dielectric tensor of crystals
EXPERIMENT
MUELLER MATRIX
JONES MATRIX
CONSTITUTIVE TENSORS
THEORY
MAXWELL EQUATIONS (Berreman formulation)
MATRIX
MULTIPLICATION of complex 4x4 matrices
(forward and backward propagating waves)
e.g. Cloude’s
General scheme of the approach:

## Slide 31: Applications. Dielectric tensor of crystals
Rutile (Uniaxial)
G. E. Jellison, F. A. Modine, and L. A. Boatner, Opt. Lett. 22, 1808 (1997).
Jellison and Baba, J. Opt. Soc. Am. 23, 468 (2006).

## Slide 32: Applications. Dielectric tensor of crystals
Rutile (Uniaxial)

## Slide 33: Monoclinic CdWO4
Jellison, McGuire, Boatner, Budai, Specht, and Singh, Phys. Rev. B 84, 195439 (2011).
Applications. Dielectric tensor of crystals
Note that a non-diagonal dielectric tensor can led to a block diagonal MM

## Slide 34: Mueller matrix Scatterometry
Measurements in periodic grating-like structures
Analysis of the Zeroth-order diffracted light (specular reflection).
Qualitative understanding of the measurements is posible attending to MM symmetries, and Rayleigh anomalies of higher orders. Energy distribution to higher orders
Trench nanostructure encountered in the manufacturing of flash memory storage cells
e-beam patterned grating structure
Expect the same symmetries as for a sample with optic axis lying in the plane of the sample
Rigorous-coupled wave analysis (RCWA). Field components expanded into Fourier series
(Form anisotropy)

## Slide 35: Mueller matrix Scatterometry
S. Liu, et al., Development of a broadband Mueller matrix ellipsometer as a powerful tool for nanostructure metrology, Thin Solid Films , in press

## Slide 36: Mueller matrix Scatterometry
S. Liu, et al., Development of a broadband Mueller matrix ellipsometer as a powerful tool for nanostructure metrology, Thin Solid Films , in press

## Slide 37: Helicoidal Bragg reflectors
Classical approximate formulas for Bragg reflection from cholesteric liquid crystals
Structural chirality, no real magneto-electric origin.
Cholesteric liquid crystal

## Slide 38: Helicoidal Bragg reflectors
Macraspis lucida
AOI 35
AOI 50
AOI 60
AOI 70

## Slide 39: Helicoidal Bragg reflectors
H. Arwin, et al. Opt. Express 23, 1951-1966 (2015).
H. Arwin et al. Opt. Express 21, 22645-22656 (2013).

## Slide 40: Plasmonic nanostructures
Typically measurements are made on 2D periodic nanostructures with characteristic dimensions comparable or smaller than the wavelength of light
Big spatial dispersion effects
d= 250 nm
a= 530 nm
The electric polarization at a certain position is determined not only by the electric field at that position, but also by the fields at its neighbors
…. And the neighbors change depending on how we orient the sample in the ellipsometer….
~

## Slide 41: Plasmonic nanostructures
RECTANGULAR
RHOMBIC
OBLIQUE
TILT
Projections of a square lattice
This what the photons of the ellipsometer will “see”
In transmission a square lattice is isotropic

## Slide 42: Plasmonic nanostructures
B. Gompf et al. Phys. Rev. Lett. 106, 185501 (2011)

## Slide 43: Plasmonic nanostructures
O. Arteaga, et al., Opt. Express., 22, 13719, (2014)
Even a highly symmetric plasmonic nanostructure must be described by a non-diagonal, asymmetric Jones matrix whenever the plane of incidence does not coincide with a mirror line

## Slide 44: I have a isotropic sample, should I study with Mueller matrix ellipsometry?
Concluding remarks
I have an anisotropic sample, can I study it with standard ellipsometry?
Yes, it never hurts. Having access to the whole MM also helps to verify the alignment of the sample.
Most likely yes, although Mueller matrix ellipsometry is arguably better suited.
Reorientations are going to be necessary. Will fail if there is some significant depolarization
Not with standard ellipsometry. Possibly with Mueller ellipsometry. But be aware! In reflection you will be NOT measuring directly optical rotatory dispersion or circular dichroism.
I have an optically active sample, can I study it with standard ellipsometry? And with Mueller ellipsometry?

## Slide 45
Summary of ideas to take home
Symmetries or assymetries of a MM give information about the orientation the sample and/or the crystallographic system
When posible (small depo) convert a experimental MM in a Mueller-Jones matrix or a Jones matrix and work from that
For intrinsic anisotropy the non-diagonal Jones elements are small and the Mueller matrix is close to a NSC matrix. If they are large suspect about structure-induced anisotropy or misalignement of the sample
Mueller matrix ellipsometry has the same applications as standard ellipsometry, plus it handles accurately anisotropy and depolarization. Important for crystals, nanotechnology, scatterometry, etc

## Slide 46
MMs at normal-incidence reflection
G. E. Jellison et al., Phys Rev. B 84, 195439(2011)
MI Alonso et al., Thin Solid Films 571, 420-425 (2014)
G. E. Jellison et al. J. Appl. Phys. 112, 063524 (2012)
O. Arteaga et al. Opt. Let. 39, 6050-6053 (2014)
MM symmetries
O. Arteaga, Thin Solid Films 571, 584-588 (2014)
H. C. van de Hulst, Light scattering by small particles, New York, Dover (1981)
R. Ossikovski, Opt. Let. 39,2330-2332 (2011).
O. Arteaga et al, Opt. Let. 35, 559-561 (2010)
J. Schellman, Chem. Rev., 87, 1359-1399 (1987)
MM scatterometry
A. De Martino et al., Proc. SPIE 6922, 69221P (2008).
S. Liu, et al., Development of a broadband Mueller matrix ellipsometer as a powerful tool for nanostructure metrology, Thin Solid Films , in press
Some further references
MMs at normal incidence transmission
DF of low symmetry crystals
MM and metamaterials
T. Oates et al., Opt. Mat. Expr. 2646, 2014.

## Slide 47
Acknowledgments
oarteaga@ub.edu
http://www.mmpolarimetry.com
R. Ossikovski (EP), A. Canillas (UB),
S. Nichols (NYU) , G. E. Jellison (ORNL)